Automotive Transient Generator

- Automotive Transient Generator Specifications

| Parameter | Specification |  |
| --- | --- | --- |
| Power Supply Details |  |  |
| Output voltage | 0V to + 60V |  |
| Output current | 0A - 50A, continuous |  |
| Peak current | 100A for max. 500ms |  |
| Frequency range | DC to 100kHz * |  |
| Supply voltage | 3x400V (US-type 3x480V) |  |
| MICROPULSES |  |
| Test voltage | U = 20V - 600V ± 10% (peak voltage and polarity as per selected standard) |  |
| Repetition rate | 0.2s - 99.0s |  |
| ISO PULSE 1 |
| Rise time | 1us, +0us/-0.5us & 3us +0us/-1.5us Please note: 10% _ 90% is settings in Oscilloscope horizontal curser for measuring Tr. |  |
| Pulse duration | 1ms and 2 ms ± 20% |  |
| Int. resistor | 10ohm and 50ohm ± 10% |  |
| ISO PULSE 2a (Need to mention about the 2b pulse) |
| Rise time | 1us +0us/-0.5us |  |
| Pulse duration | 50us ± 20% |  |
| Int. resistor | 2Ω ± 10% |  |
| ISO PULSE 2b |  |
| Rise time | 1ms +/-0.5ms |  |
| Pulse duration | 0.2s to 2s |  |
| Int. resistor | 0Ω to 0.05Ω |  |
|  | As per ISO 7637-2 standard we need cover Pulse 2B also |  |
| ISO PULSE 3a (test requirement should meet as per ISO 7637-2 pulse 3a and 3b) |
| Rise time | 5ns +/- 1.5ns It is 20% only, if they are no mentioned any tolarances.since they given tolarances no need 20%. As per 7637-2_2004 Standard P.g no 13. |  |
| Pulse duration | 0.1us, +0.1us\ -0us _as per 2004 standard And 150ns +\- 45ns as per 2011. Above comment is applicable |  |
| Int. resistor | 50 ohm ± 10% |  |
| Burst duration | t4= 10ms |  |
| Burst repetition rate | t5 =90ms |  |
| Spike frequency | f = 0.1kHz - 200kHz t1 = 5μs - 10ms |  |
| ISO PULSE 3b |
| Rise time | 5ns +/- 1.5ns |  |
| Pulse duration | 0.1us, +0.1us\ -0us _as per 2004 standard And 150ns +\- 45ns as per 2011. |  |
| Int. resistor | 50 ohm ± 10% |  |
| Voltage (Open circuit ) | 25V – 1000V ± 10 % |  |
| At 50 Ohms | U = 13V – 500V ± 10 % |  |
| Rise time | 5ns +/- 1.5ns |  |
| Burst duration | t4= 10ms |  |
| Burst repetition rate | t5 =90m |  |
| Spike frequency | f = 0.1kHz - 200kHz t1 = 5μs - 10ms |  |
| ISO PULSE 4 |
| tf [Vb-Va1] | < 5ms |  |
| Vb | 0.0V – 30.0V (60.0V) |  |
| Va1 | -30.0V(-60.0V)- 30.0V(+60.0V) |  |
| Va2 | -30.0V(-60.0V)- 30.0V(+60.0V) |  |
| t1 | 0.1S – 99.9S |  |
| t7 | 15ms to 100ms |  |
| t8 | 5ms to 50ms |  |
| t9 | 0.1s to 20 s |  |
| t11 | 5ms to 100ms |  |
| Va | 0V-60V |  |
| ISO PULSE 5a and 5b (test requirement should meet as per ISO 7637-2 and 16750-2) |
| Rise time | 10ms,-5ms\+0ms |  |
| Pulse duration | 400ms |  |
| Int. resistor | 0.5 ohm to 8 ohm ± 10% |  |

- Conducted Transient Emission

| Parameter | Specification |  |
| --- | --- | --- |
| Test voltage |  |
| Operating voltage | Max. 60 V |  |
| Operating current | Max. 100 A continuous |  |
| Peak current protection | 500 A |  |
| Inrush current | 400 A for 200ms |  |
| Voltage drops | Less than 1V @ 25 A Umax=400V @25A Imax 25 A continuously,100 A for t ≤ 1 s. switching time, Δts 300 ns 20 % with DUT; |  |
| Overvoltage protection | By Varistor |  |
| Overload protection | Short – circuit over temperature protected switch off after approx. 2 min with 120 A; switch on after approx. 45 s cooling time. |  |
| Inverse polarity protection | Protected with an additional acoustic signal in case of inverse polarity |  |

- Electronic Switch

| Parameter | Specification |  |
| --- | --- | --- |
| Test voltage |  |
| Switching time | 300 ns ±20% (240 ns – 360 ns) into test load 50µH/0.6Ω |  |
| On/Off duration | Min. 10ms to 500ms continuously selectable by potentiometer ± 5% |  |
| On/Off repetition | Min. 1 s to 10 s continuously selectable by potentiometer ± 5% |  |
| Operation | Switch closed indicated by LED |  |
| Trigger |  |
| Auto | Automatic trigger with min. ~ 0.1 Hz to max. 1 Hz repetition, continuously selectable by potentiometer |  |
| Extern | External trigger ↓0V, BNC input (Umax. +15 V) |  |
| External LV124 Mode | Active when Manual Trigger button is pressed during “Power ON” |  |
| Trigger delay typical | Switch off: approx. 10.5 us Switch on: approx. 94ms |  |
| Manual | Manual trigger of a single event |  |
| Requirement of Power Supply adapter : It is necessary if, external supply is required to turn on the switch |  |

- Single Line Artificial Network

| DUT Supply | Specification |  |
| --- | --- | --- |
| Max operating voltage | 1000V DC / 250V AC (up to 1kHz) |  |
| Max. operating current | 50A continuous |  |
| Max. peak current | 100A |  |
| Max operating voltage | 1000V DC / 250V AC (up to 1kHz) |  |
| Artificial Network | Specification |  |
| Impedance | 50 Ω║ 5 μH + 1 Ω +/- 10% |  |
| Frequency range | 100 kHz to 108 MHz |  |
| Inductance | 5 μH (Air core) |  |
| Insertion loss | < 3 dB DUT to receiver output |  |
| Coupling capacitor | 0,1μF |  |
| Connector | Multi Contact 6mm High-current connector up to 100A 4mm banana safety lab connector up tor 32A |  |
| BNC plug | for 50Ω terminating resistor or measuring device (in parallel with 22kΩ) |  |

Electrical transient transmission by capacitive and inductive coupling via lines other than supply lines

- capacitive coupling clamp (CCC) method:

- Direct capacitive coupling (DCC) method:

- Inductive coupling clamp (ICC) method:

Additional Items with Automotive Transients (But not limited to):

- Calibration kit and Calibration Fixture for Transients:

- Software for equipment operation:

- Batteries 6 numbers (4 numbers of 12V DC,90AH. 2 numbers of 12V DC 150AH):

- Battery Charger as required to charge above batteries:

- BNC cables (As required for using with test instrument):

- Insulation Support _50mm thickness(Length = 1.2m width :0.5m (Styrofoam):

- Capacitive Coupling Clamp (CCC method) (1 number):

- Injection Probe (ICC method) along with calibration jig and accessories:

- DCC method Capacitor value 100pF and 0.1uf and 470pF:

- Two numbers of High Voltage LISN acc. To CISPR 25 Ed. 4 or BMW GS 95025-1 to measure the conducted disturbance voltage on shielded lines for (hybrid) electric vehicles (HEV, EV), can be used for BCI with an external dummy load, impedance (5μH) || 50 Ohm. 70 (100) A, 1000 V DC:

- Shielded housing for 2 HV-LISN, 2 paths with cable feed troughs for HV+ and HV-, shield can be connected to the housing, 2 measurement ports N, 2 monitor ports N, with connecting cables between inside measurement ports and outside N-connectors:

| Description | Specification | Quantity |
| --- | --- | --- |
| Metallic Table | Table of size H:0.9mxW:1.2mxL:2m | 3 |
| Ground reference plane | 2.44m (Length) X 4m (Width) min 0.25mm to 0.65mm thickness (Aluminum or copper) - Checked Plates | 3 |
| Ground reference plane | 2.0m (Length) X 4m (Width) min 0.25mm to 0.65mm thickness (Aluminum or copper) - Checked Plates | 1 |
| MULTIMETER | Voltage up to 1000V AC and DC; Current rating up to 10A, with Acc. calibration certificate. | 1 |
| CLAMP METER | Voltage up to 1000V AC and DC; Current rating up to 100A, with Acc. calibration certificate. | 1 |
| RC Network | Capacitor = 220pF ± 20%, Resistor 510Ω ± 10% | 1 |
| Test Rack | Net-rack Table 1.5mL x 0.7mWx2mH with accessories | 2 |
| Insulation Support | Styrofoam Table (0.8mX1.0mX0.8m) (H*L*W) | 1 |
| Wooden Insulation Support | Wooden Support, (0.1m*1.5m*1.2m) (H*W*L) Wooden Support, (0.1m*0.6m*0.6m) (H*W*L) | 1 each |
| ESD Target | Target as per IEC 61000-4-2 (Only target & Fixture not required) | 1 |
| ESD Target accessories | Calibration adopter, 20db SMA Attenuator (additionally), SMA cable, SMA to BNC adopter (For cable chain) | As required for target |
| Oscilloscope | 2GHz Bandwidth | 1 |

Supported Standards

ISO 7637-1

ISO 7637-2

ISO 7637-3

ISO 16750-2

AIS 004 part 3

ECE R10 Rev 5